Procedure: Filtering analysis for Particulate Inorganic Carbon (PIC)

Other names: CaCO3; particulate Ca

Date: 9 May 2024

Author: Alex Poulton

Key References/Reading: Poulton et al., 2006 GBC; Daniels et al., 2012 L&O; Daniels, 2010 (unpublished lit review on different methods for determining CaCO3 concentrations)

Protocol:

- Sampling: Water samples (0.5-1 L) collected from CTD sampling depths or underway seawater supply into brown HDPE bottles (Nalgene) via silicon tubing (Nalgene). For surface samples from the CTD and underway seawater supply, triplicate samples are collected and processed.

- 0.5 L measured out from each depth and filtered through 0.45 µm 25 mm dia. polycarbonate (e.g. NucleoporeTM) filters under low pressure.

- Once sample has passed through the filter, rim and cup rinsed extensively with pH adjusted MilliQ (see notes) to ensure all particulate material collected on filter and to remove any remaining saltwater residue. Multiple rinses of the filter with pH adjusted MilliQ recommended.

- Filters removed from filtration fit and placed in 15-mL Falcon tubes, suitably labelled, and placed in an oven (40-50oC) over night (~12 hrs) to dry. Tubes are left without lids at this point, so tube needs to be labelled. After drying tubes capped and placed as sample set in Ziploc bag labelled with CTD or event number and kept in a dry and safe place before analysis.

- Analysis: Samples are extracted 0.5 mL of 0.4 M high-purity nitric acid and analysed using Inductively Coupled Plasma Atomic Mass Spectrometry (ICP-OES) at the University of Maine.

Notes:

- Previous protocols have used 0.02 M potassium tetraborate solution (e.g., Poulton et al., 2006) to wash saltwater from the filters; the key thing to consider is to remove saltwater (NaCl) which contains large amounts of dissolved Ca (~10 mmol L-1) that would swamp the particulate Ca signal. The Na signal is also used to assess saltwater contamination and due to the high concentration of K in seawater (~10 mmol L-1) the protocol has moved to use pH adjusted MilliQ (after Daniels et al., 2012).

- The pH of MilliQ is typically slightly acidic (~7) so the pH is adjusted by adding a trace amount (180 µL) of Ammonia solution (Ammonium hydroxide, 25%) to adjust the pH to around 9-10.

- The pore size of the filters used in 0.45 µm which should collect all particulate CaCO3.