Analysis of PIC Filters for (Ca, Mg, Sr, Na)

at Bigelow Laboratory for Ocean Sciences

(UMaine procedures from 09-25-24)

Leaching of filters:

- Filters were mailed to us in 15 mL centrifuge tubes

- Add 5 mL of ultraclean 5% nitric acid to each tube (Fisher Optima grade)

  - Cap, shake, and make sure filter is submerged in acid

- Let sit for ~24 hours (overnight)

- Samples are ready to be run

  - Samples are transferred to a polypropylene 4mL vials that have been acid cleaned.

Standards:

- We made mixed standards of Ca, Mg, Sr and Na out of mono-elemental solutions, at various concentrations of each element, for our calibration curves

  - All standards were mixed in ultraclean 5% nitric acid

  - All standards contained 1 ppm of K (not analyzed)

  - Standard curves are made from a mixed standard of Ca, Mg, and Na and a separate standard is used for strontium.

  - Calibration curve notes:

      - for Ca, Mg and Na – calibration curves typically used the standards from 10 to 5000 ppb and Sr from 1 to 100ppb.

  - Additionally, to test instrument stability and reproducibility, a standard is run approximately every 10 – 20 samples.  To verify the calibration a certified reference material is run once for each run.

Instrument Parameters:

- All samples were analyzed on a Thermo Element XR HR-ICP-MS, 87Sr was measured in low resolution, 43Ca, 26Mg, and 23Na were measured in medium resolution.

- A blank is included in each calibration curve no additional blank subtraction is done.

- Detection limits calculated September 2024

  - Mg - 0.5 ppb

  - Ca - 2.5 ppb

  - Sr - 0.08 ppb

  - Na – 4.0 ppb